## Slide 1
Order Definition
Order Execution
PCA MSIS
Patient
ICS Boundary
PoC Process Boundary
Order Governance Boundary
ICS Component Boundary
HDO Organizational Boundary
Order Construction
HDO
Policies
Government / Public Policy Context
Control
Feedback

## Slide 2
2800-2-1
2800-3-1-1
2800-1-1
Mock Vendor
Submission
2800-0
Requirements on how to fill in the gaps (document/disclosures) in “templates” for an ICE-based PCA Monitoring ICS
Requirements on what to account for in Risk Management
Requirements on what to test and how to document results
Requirements on structure/content of final assurance case
Implementation
Compliance
ICE Architecture Platform and Service Component Implementation
Uses
App
System/App Requirements Document
Risk Management Results
Verification Results
Assurance Case
Compliance
To Clinical and safety Objectives
Compliance to architecture and development requirements
Implementation
Compliance
2800-1
Compliance to architecture specification requirements
Compliance
to Requirements to map to Safety, Security, and Essential Performance Objectives
2800-2

## Slide 3: Boundary Definition Activities
Delineate Boundary
Identify inputs/outputs (directional modes)
Identify Actor Categories (human/machine) on either side of interface
From Component’s point of view
  Inputs (flow from context to component)
    specify assumptions ACmp on inputs (requirements on context)
  Outputs (flow from component to context)
    specify guarantees GCmp on outputs (requirements on component)
  Acmp and GCmp form the Component’s interface contract
From Context’s (Client’s) point of view
  Call/Push (flow from context to component)
    Specify guarantees GCxt on flows to component (requirements on context)
  Return/Pull (flow from component to context)
    Specify assumptions ACxt on flows from component (requirements on component)

## Slide 4: Assembly/Integration Assurance
Safety case will include the following assurance steps for parameters on which SSEP depends
  GCxt complies with ACmp
  GCmp complies with ACxt
SSEPOs will state the “vocabulary” or categories of 2800-relevant properties mentioned in AX, GX
Testing/Verification activities will provide evidence of the compliance steps
Component
Context
ACmp
ACxt
GCmp
GCxt

## Slide 5: Implementation Compliance Assurance
Safety case will include the following assurance steps for parameters on which SSEP depends
  The component implementation which includes internal components providing properties Pi complies with its interface contract (Acmp,GCmp)
SSEPOs will state the “vocabulary” or categories of 2800-relevant properties mentioned in AX, GX
Testing/Verification activities will provide evidence of the compliance steps
Note that the ICS (System) interface is a special instance of the component interface
Component
Interface
ACmp
GCmp
P2
P1
P3
Component
Implementation
Where do we get, e.g., the system level A,G related to safety? Ans: aim to ensure safety (avoid hazards) related to controlled process (patient).

## Slide 6
Controller Setting
Operator
Actuation
Via SUI App Control Panel
Control Actions
Feedback, sensed properties
Generic Control Loop:
Controller
Actuator
Sensor
Controlled Process
Model of Controlled Process
Key concept: Avoid unsafe control actions (i.e., control actions leading to hazards). Safety-related requirements are those that are needed to avoid unsafe control actions.

## Slide 7
Control
Action
Generation
Model of
Automation
Model of
Controlled Process
Human Controller
Control
Action
Generation
Model of
Controlled Process
Automated Controller
Controlled Process
Written/Trained
Procedures
Environmental
Inputs
Process Inputs
Process Outputs
Disturbances
Actuators
Sensors
Controls
Displays
Stacking Control Actions – combining human controller (e.g., clinician) that controls and automated controller (e.g., ICS)

## Slide 9
Role: Patient
Role: Clinician (App Operator)
App Boundary
SUI App Display
Role: Data Logger
PCA
Pump
M.x
System (ICS) Boundary
Role:
Device
PoC Process Boundary
E.x
E[X]
Role: External System
Role: App
Role: ICS
Monitoring Devices
Role:
Device
AppDisplayToOper
OperToAppDisplay
ICEDisplayToOper
OperToICEDisplay
M[X]ToOper
OperToM[X]
PatientToM[X]
PumpToOper
OperToPump
PumpToPatient
PatientToPump
This diagram is an abstraction of the 10746 Computation View
AppToM[X]
M[XToApp]
PumpToApp
AppToPump
AppToLogger
OrderGovernanceToPoC
PoCToOrderGovernance
AppToEx.X
Ex[X]ToApp
ICE Platform Boundary
Pump To Ex.X
Ex[X ]To Pump

## Slide 10
Patient
SPO2
(Pulse Ox)
RR
(capnography)
Clinician (App Operator)
App Boundary
SUI App Display
ICE Data Logger
PCA
Pump
EtCO2
(capnography)
System (ICE) Boundary
Device Suite
Process Boundary
a
b
c
d
e
f
g
h
i
j
k
l
n
m
o
p

## Slide 11
Patient
SPO2
(Pulse Ox)
RR
(capnography)
Clinician (App Operator)
App Boundary
SUI App Display
ICE Data Logger
PCA
Pump
EtCO2
(capnography)
System (ICE) Boundary
Device Suite
Process Boundary
a
b
c
d
e
f
g
h
i
j
k
l
n
m
o
p
1
2*
3
4
5

## Slide 12
Operator
Cognition
Operator
Actuation
Via SUI App Control Panel
Operator
Sensing
Via SUI App Display
Application
Control
Algorithm
Parameters (set points) for control algorithm
Algorithm and patient status via App Supervisor Display
Control Loop 1: Clinician-AppControl
Clinician

## Slide 13
Operator
Cognition
Operator
Actuation
Via Device Control Panel
Operator
Sensing
Via Device Display
Device Control
Parameters for Infusion Prescription
Infusion Status,
Technical Alarms
Control Loop 2a: Clinician-Device-PCAPump
Clinician

## Slide 14
Operator
Cognition
Operator
Actuation
Via Device Control Panel
Operator
Sensing
Via Device Display
Device Control
Parameters for Limit Alarm Settings, Device Actuation
Physiological Parameter Values,
Technical and Limit Alarm Notification
Control Loop 2b: Clinician-Device-Capnography
Clinician

## Slide 15
Operator
Cognition
Operator
Actuation
Via Device Control Panel
Operator
Sensing
Via Device Display
Device Control
Parameters for Limit Alarm Settings, Device Actuation
Physiological Parameter Values,
Technical and Limit Alarm Notification
Control Loop 2c: Clinician-Device-PulseOximeter
Clinician

## Slide 16
Operator
Cognition
Operator
Actuation
Via Voice, Written Instructions
Operator
Sensing
Sight, Sound
Patient
Patient Instructions for PCA Pump operation and care giving procedures
Visual and verbal cues from patient regarding pain level and physical condition
Control Loop 3: Clinician-Patient
Clinician

## Slide 17
Patient
Cognition
Patient
Actuation
Via Pump Bolus Button
Patient
Sensing
Sight, Sound
PCA Pump
Button Press
Audio/Visual cue of button press or infusion mode displayed on device front panel?
Control Loop 4: Patient-BolusDose
Sight/Sound recognition of device front panel cues and pain relief sensation
Patient

## Slide 18
App
Control
Algorithm
PCA Pump Actuation
RR
Device
Sensing
Patient
Flow of Opioid
Physiological condition
Control Loop 5: App-Device-Loop
Physiological parameter readings,
Limit alarm notifications
Technical alarm notifications
Pump KVO command
Pump Normal command
SpO2
Device
Sensing
EtCO2
Device
Sensing

## Slide 19
App
Control
Algorithm
PCA Pump Actuation
RR
Device
Sensing
Patient
Flow of Opioid
Physiological condition
Control Loop 5: App-Device-Loop
Physiological parameter readings,
Limit alarm notifications
Technical alarm notifications
Pump KVO command
Pump Normal command
(from general set of pump management and flow control commands)
SpO2
Device
Sensing
EtCO2
Device
Sensing
(with classifications of communication links and component types)
[Link Role: Physical substance continuous flow,
Medium Risk Responsibility: Pump Vendor,
Content Risk Responsibility: Pump Vendor]
[Component Role: Physical substance continuous flow actuation (infusion device),
Risk Responsibility: Pump Vendor]
[Link Role: Actuator Control,
Medium Risk Responsibility: Network Controller Provider,
Content Risk Responsibility: App provider]
[Link Role: Physiological Data,
Medium Risk Responsibility: Network Controller Provider,
Content Risk Responsibility: Data Stream Provider]
[Component Role: Physiological Parameter Sensor (Transducer),
Risk Responsibility: Sensing Device Vendor]
[Link Role: Physiological Parameter Sensor and associated communication,
Risk Responsibility: Sensing Device Vendor]
[Component Role: Controller – Interlock,
Risk Responsibility: App Vendor]

## Slide 20
Physician
Cognition – Formation of Order
Recording and
Storage of Order
Recording of patient’s health in EMR, etc.
Order Execution Process
(Clinical Process?)
Transmission of Order to Care-giving Context
Outcomes of Order Execution reflected in patient’s actual health
Control Loop: Order Definition
Mental Model of Patient’s Health

## Slide 21
Interpretation of Order by Clinicians
Actions to Configure ICS PCA, Patient
Output of ICS PCA
Direct Observations of Patient’s Health
ICS PCA State
Patient State
Control Loop: Clinical Process
Mental Model of Patient’s Health and ICS PCA State
Output to Physician

## Slide 22
Configuration
Self Test
Operational
<operator
requests
re-configuration>
<operator requests
operational mode>
<self test
succeeds>
<self test
fails>

## Slide 23
Configuration
Self Test
Operational
<operator
requests
re-configuration>
<pre-conditions satisfied &
operator requests
operational mode>
<self test
succeeds, fails>
<operator
requests
self-test>

## Slide 24
Patient
SPO2
(Pulse Ox)
RR
(capnography)
Clinician (App Operator)
App Boundary
SUI App Display
ICE Data Logger
PCA
Pump
EtCO2
(capnography)
System (ICE) Boundary
Device Suite
Process Boundary
a
b
c
d
e
f
g
h
i
j
k
l
n
m
o
p
5
4 variable model notes
Add communication from pump to app

## Slide 25
Role: Patient
Role: Clinician (App Operator)
App Boundary
SUI App Display
Role: Data Logger
PCA
Pump
M.x
System (ICS) Boundary
Role:
Device
PoC Process Boundary
E.x
E[X]
Role: External System
Role: App
Role: ICS
Monitoring Devices
Role:
Device
AppDisplayToOper
OperToAppDisplay
ICEDisplayToOper
OperToICEDisplay
M[X]ToOper
OperToM[X]
PatientToM[X]
PumpToOper
OperToPump
PumpToPatient
PatientToPump
This diagram is an abstraction of the 10746 Computation View
AppToM[X]
M[XToApp]
PumpToApp
AppToPump
AppToLogger
OrderGovernanceToPoC
PoCToOrderGovernance
AppToEx.X
Ex[X]ToApp
ICE Platform Boundary
Pump To Ex.X
Ex[X ]To Pump

## Slide 26
Patient
SPO2
(Pulse Ox)
RR
(capnography)
Clinician (App Operator)
App Boundary
SUI App Display
ICE Data Logger
PCA
Pump
EtCO2
(capnography)
System (ICE) Boundary
Device Suite
Process Boundary
a
b
C mp_spo2)
e
f
g
C (m_spo2)
j
k
l
n
m
o
p
4 variable model notes
Add communication from pump to app
System Monitored Variables
C mp_spo2)
Latency,jitter
Sensing delay, processing, resolution, etc.
System Timing Requirements
A Cp_shutoff
A Cp_shutoff
Flow Rate
Command
Shutoff Delay.

## Slide 27
Device Output
M to Mp – constraints on device, exposed as app requirements (formal specs) on device interface
Mp to Mp-a – communication across the network from device output to app input

## Slide 28: 4 Variable Model Notes
Consider the control loop on the previous slides that involves physio parameter sensors, app-based “controller”, and PCA pump as actuator
For the system property capturing the “shutoff requirement”, introduce 4-var naming scheme for different system boundary inputs and outputs and component inputs/outputs for monitoring devices, app, and PCA
Imagine a simple shutoff condition based on some monitoring device parameter going out of range. Using the 4-variable paradigm, state the requirement.